SISTEM  ZA  ELEKTROHEMIJSKA  ISPITIVANJA

NA  BAZI  PC  I  LABVIEW  PAKETA

SYSTEM FOR ELECTROCHEMICAL INVESTIGATIONS

BASED ON PC AND LABVIEW PACKAGE

Zoran Stević, Mirjana Rajčić-Vujasinović, Aleksandra Milkovski

Univerzitet u Beogradu, Tehnički fakultet u Boru

Sadržaj - U radu je opisan sistem za elektrohemijska ispitivanja baziran na PC računaru i softverskom paketu LabVIEW. Dat je pregled poznatih elektrohemijskih metoda i parametara, počev od merenja potencijala i najjednostavnijih (hronopotenciometrija i hronoamperometrija) pa do elektrohemijske impedantne spektroskopije. Poslednja naznačena metoda prilagođena je merenju na sistemima koji sadrže vrlo visoke kapacitivnosti, što je postalo aktuelno tek sa pojavom elektrohemijskih superkondenzatora.

Abstract – A system for electrochemical investigations based  on PC computer and software package LabVIEW was developed and described in the paper. It was given an overview of well known electrochemical methods and parameters, beginning with potential measurement and simple methods such as chronopotentiometry and chronoampermetry, till electrochemical impendance spectroscopy. The last named method is adapted for systems containing large capacitancies, that became actual with appearance of electrochemical supercapacitors.

1. UVOD

Postoji niz standardnih metoda za ispitivanje elektrohemijskih sistema [1-13]. Ovde će biti opisana realizacija najvažnijih standardnih metoda na sistemu baziranom na PC računaru i softverskoj platformi LabVIEW.

Merno-upravljački interfejs i softver koji su rezultat  sopstvenog razvoja u ovom trenutku omogućavaju primenu sledećih metoda:

- merenje mirujućeg potencijala

- galvanostatska metoda

- potenciostatska metoda

- ciklična voltametrija

- hronoamperometrija

- hronovoltametrija

- pulsne tehnike

- naponsko svipovanje

- strujno svipovanje

- impedantna metoda

Posebno treba istaći elektrohemijsku impedantnu spektroskopiju (EIS), kod koje su postignute performanse koje se analognom tehnikom nisu mogle postići (donja granična učestanost je 1Hz). Kod elektrohemijskih sistema koji uključuju velike kapacitivnosti (kao što su superkondenzatori npr.) neophodne su tako niske učestanosti ispitnog napona. Takvi sistemi ispituju se u novije vreme, a čak i da je ranije postojala potreba za merenjem pri ovako niskim frekvencijama, to nije bilo moguće uspešno sprovesti bez najsavremenije računarske tehnike koja je uključena u opisani sistem.

2.  HARDVER

Za generisanje signala i praćenje odziva ispitivanog elektrohemijskog sistema razvijen je merno-upravljački sistem na bazi PC Pentijum računara. Hardverski deo pored računara obuhvata komercijalni AD-DA konvertor i spoljašnji interfejs za obradu analognih signala i oni će biti posebno opisani [14-23].

Za ADDA konverziju upotrebljen je komercijalno dostupan konvertor Burr-Brown PCI-20428 W [24]. On sadrži 16 analognih ulaznih kanala sa 12-bitnom konverzijom i dva analogna izlazna kanala, takođe sa 12-bitnom konverzijom.

Merni interfejs je sklop koji je za potrebe naponskih i strujnih ispitivanja elektrohemijskih sistema bilo  neophodno dizajnirati sa sledećim karakteristikama:

- dva upravljačka naponska ulaza  ± 10 V

- jedan merni strujni ulaz  ± 20 mA

- jedan naponski izlaz  ± 10 V za ulaznu struju ± 20 mA

- jedan naponski izlaz za praćenje referentnog potencijala

- otpornost ulaza za referentnu elektrodu veća od 1012 Ω

- jedan troelektrodni izlaz za elektrohemijsku ćeliju sa ručnim izborom:

- naponski u opsegu ± 1 V sa mogućnošću superponiranja malog signala u opsegu ± 10 mV

- strujni u opsegu ± 10 mA

Navedeni zahtevi ostvareni su sklopom čija je blok šema data na slici 1.

Slika .1.  – Blok šema interfejsa

Zadati oblik naponskog signala dovodi se sa analognog izlaza – kanal 0 (AO - CH0) na oslabljivač A2 i posle deljenja sa 10 sabira sa drugim upravljačkim naponom (AO – CH1) oslabljenim 1000 puta, tako da se na izlazu sabirača A3 formira Uref = -V1 – V2/1000 = -VCH0 / 10 – VCH1/1000 čime se postiže mogućnost  izlaza ± 1 V sa superponiranim ± 10 mV (za upravljačke napone ± 10V).  Izlazni pojačavač – komparator K postavlja taj isti napon na svom izlazu u odnosu na referentnu elektrodu (RE). Radna elektroda (WE) je praktično uzemljena (preko mernog otpornika Rš), pa se zato kontraelektroda (CE) negativno polariše kada se želi pozitivna polarizacija radne u odnosu na kontra elektrodu.

Zaštitno kolo (ZK) prati napon na AO-CH0 i drži isključen kontakt relea (P2) do inicijalizacije AD-DA ploče i uspostavljanja radnog režima.

Za slučaj strujnog režima predviđen je konvertor (U/I) čiji odnos 1mA/1V obezbeđuje izlaznu struju ± 10 mA za upravljački napon ± 10V. Izbor režima vrši se ručno preklopnikom P1.

Podsklopovi opisane blok šeme standardno su izvedeni, uz posebnu pažnju posvećenu ulaznoj otpornosti priključka RE.

Za postizanje velike ulazne otpornosti pojačavači A1 i K posebno su izvedeni na jednoj odvojenoj pločici, a upotrebljeni su operacioni pojačavači sa velikom ulaznom otpornošću.

- SOFTVER

Platforma za pisanje aplikacija za predviđene merne metode bio je paket LabVIEW 6.1 (National Instruments) [25-30], koji važi za visoki standard u svetu merne tehnike i virtuelnih instrumenata. Za vezu sa AD-DA konvertorom, odnosno spoljašnjim svetom korišćena je Master Link kolekcija drajvera koja se dobija uz konvertor PC1 – 2048W [25].

U daljem tekstu biće opisane aplikacije za neke od postavljenih elektrohemijskih metoda.

GALVANOSTATSKA METODA

Šema aplikacije za ovu metodu u LabVIEW paketu data je na slici 2.

Slika 2.- Dijagram aplikacije za galvanostatsku metodu

U delu inicijalizacije upotrebljeni su sledeći blokovi :

Init SW – softverska inicijalizacija

428W – inicijalizacija ploče

AssignIO – naznačenje ulaza – izlaza

HW Init – hardverska inicijalizacija

Slot – izbor slota adrese ploče

Address – izbor adrese u okviru slota

AO Conf – konfiguracija izlaza

Channel – broj kanala

Module – broj modula

Range – opseg napona (= 4 za ± 10 V)

Očitavanje napona vrši se na kanalu AI CH0 i posle usrednjavanja, korigovanja i umanjenja za uneti mirujući napon (Emir [V])  dobija se nadnapon kao odziv na galvanostatsku ekscitaciju. On se prikazuje na indikatoru eta [V]  i na dijagramu u funkciji vremena.

Intenzitet struje zadaje se naponom na izlazu bloka AOWrite u odnosu 1 V / 1 mA. Konvertor napona u struju hardverski je rešen u okviru spoljašnjeg interfejsa.

Trajanje galvanostatskog impulsa zadaje se kontrolom tp [s], poredi sa aktuelnim vremenom i IF blokom definiše kraj impulsa. Merenje relaksacionog nadnapona se nastavlja do dostizanja zadatog ukupnog vremena tu [s], kako je to regulisano komparatorima i uslovom za zaustavljanje WHILE petlje.

Na slici 3 prikazan je kontrolni panel sa kontrolama za zadavanje struje  I [mA], trajanja pulsa tp [s], mirujućeg potencijala Emir [V], ukupnog vremena tu [s], kao i indikatorima aktuelne struje, nadnapona i vremena. Na panelu je još i grafički prikaz           E = f(t).

Slika 3 –  Kontrolni panel za galvanostatsku metodu

POTENCIOSTATSKA METODA

Na slici 4 prikazan je kontrolni panel sa  kontrolama za zadavanje intenziteta i trajanja naponskog impulsa, mirujućeg potencijala i ukupnog vremena, kao i indikatorima aktuelnih veličina, uz grafički prikaz  i = f(t).

Na slici 5 prikazana je šema aplikacije za potenciostatsku metodu. Pošto se radi sa dužim vremenima može se ići na jedno očitavanje u sekundi što dozvoljava veće usrednjvanje a time i tačnije merenje. Ovde se trajanje pulsa definiše zadatim tp i IF blokom, a ukupno vreme eksperimenta zadatim tu i uslovom zaustavljanja WHILE petlje.

Kod svih naponskih metoda  je očitavanje struje na prvom kanalu, a zadavanje napona na nultom izlaznom kanalu.

Slika 4 –  Kontrolni panel za potenciostatsku  metodu

Slika 5 - Dijagram aplikacije za potenciostatsku metodu

LITERATURA

[1]. F. Scholz, Electroanalytical  methods: guid to experiments and application, Web reference: http://www.chemweb.com/arforms?fid=31

[2]. Th. Christen, Ch. Ohler, Optimizing energy storage devices using Ragone plots,  Journal of Power Sources  110, 1 (2002) 107-116

[3]. S. A. Brazill, Sh. E. Bender, N.E. Hebert,  J.K.  Cullison,   E.W. Kristensen   and W.G. Kuhr,   Sinusoidal   voltammetry:   a  frequency   based   electrochemical          detection technique, Journal of Electroanalytical Chemistry  531, 2 (2002)  119-132

[4]. A. Chu, P. Braatz, Comparison of commercial supercapacitors and high-power lithium-ion batteries for power-assist applications in hybrid electric vehicles           I.Initial characterization, Journal of Power Sources 112 (2002)  236-246

[5]. I.L. Skryabin, G. Evans,  D. Frost, G. Vogelman, J.M. Bell, Testing and control issues in large area electrochromic films and devices,  Electrochimica Acta 44 (1999) 3203-3209

[6]. P.C.Butler, J.F. Cole, P.A. Taylor, Test profiles for stationary energy-storage  applications, Journal of Power Sources 78, 1-2 (1999) 176-181

[7]. A.D. Pasquier, I. Plitz, J. Gural, S.Menocal and G. Amatucci, Characteristics and performance of 500 F asymmetric hybrid  advanced supercapacitor   prototypes,         Journal of Power Sources   113, 1 (2003)  62-71

[8]. Конденсаторы. Методы   измерения   электрических параметров. Общие положения. ГОСТ 21315.0-75, Издательство стандартов, 1976

[9]. C. Emmenegger, P. Mauron,   P. Sudan,   P. Wenger,  V. Hermann,   R. Gallay,  A.Zuttel,   Investigation  of  electrochemicall double-layer (ECDL) capacitors         electrodes based on carbon nanotubes and activated carbon materials, Journal of Power Sources, Web reference: http://www.elsevier.com/locate/jpowsour

[10]. N.Khan,   N. Mariun, M. Zaki, L. Dinesh,   Transient analysis of pulsed charging in  supercapacitors,  TENCON  2000,  Kuala  Lumpur,   Malaysia,   September,    2000,         Proceedings 3, 193-199

[11]. Y. Guo, J. Qi, Y.  Jiang, S.   Yang, Z.  Wang, H. Xu,  Performance  of   electrical double   layer  capacitors  with porous   cabons derived from rice husk, Materials         Chemistry and Physics 80, 3 (2003) 704-709

[12]. G. P. Dai,   M. Liu,    D.M.   Chen,     P. X.   Hou, Y.   Tong, H. M.   Cheng, Electrochemical charge-discharge capacity of purified single-walled carbon nanotubes, Electrochemical  and Solid-State Letters 5, 4 (2002) E13-E15

[13]. M. Dolle, S. Grugeon, B. Beaudoin, L. Dupont, J.-M. Tarascon, In situ TEM study of the interface carbon/ electrolyte, Journal of Power Sources  97-98 (2001) 104-106

[14] B.E.Conway, Electrochemical supercapacitors, Kluwer Academic/Plenum Publishers, New York, 1999

[15]   R. Kotz, M. Carlen, Principles and applications of electrochemical capacitors, Electrochimica Acta 45 (2000) 2483-2498

[16] G. Clerici, The supercapacitor, Tecnologie Elettriche  20, 6 (1993) 64-67

[17]  A. Bindra, Supercapacitor powers new applications, EE Times  923(1996) 1-2

[18]  R.J. Bozeman, System of memorizing maximum values,  US Patent No: 5539402

[19]  M. Rajčić-Vujasinović,      Z. Stanković,  Z.   Stević,  The   consideration  of    the electrical   circuit  analogous  to   the copper  or  coppersulfide/electrolyte interfaces   based on  the  time   transient  analysis,    Elektrokhimiya 35,  3 (1999) 347- 354

[20] D. Firsich, Carbon supercapacitor electrode materials, Patent: Inorganic Specialist Inc.  U S  Patent No: 5993996

[21]   M. Francis, Supercapacitor electrode and method of fabrication thereof – active electrode coated with carbon particles; laminated to dielectric separator, Patent: Motorola Inc. US Patent No: 5079674

[22] H. Wojtek, Process for manufacturing supercapacitor, Patent: Yardney Technical Products Inc.  US Patent No: 5649982

[23]   X. Andrieu, L. Josset, Bifunctional electrode for an electrochemical cell or a supercapacitor and a method of producing it. Patent: Accumulateurs Fixes et de Traction S A SAFT FR, US Patent No: 5811205

[24] R. Carlin, H. De Long, Supercapacitor electrochemical cell – reversible cell with nickel or palladium cathode which can oxidize to thin chloride film, for delivering high current density charge. Patent: US of America Air Force Secretary, US Patent No: 5585999

[25] BUR-BROWN   COMPANY,   Master  Link Software  Libraries,Intelligent Instrumentation Inc., Tucson, USA, 2000

[26] National instruments, LabVIEW development guidelines, NI Corporation, 2000

[27] National instruments,Integrating the internet into your measurement system, DataSocket technical overview, NI Corporation, 2000

[28] National instruments, LabVIEW analysis concepts, NI Corporation, 2000

[29] National instruments, LabVIEW getting started with Labview, NI Corporation,  2000

[30] National instruments, LabVIEW measurements manual, NI Corporation, 2000